Technical Specification – Data Sheet

# Water Purification System

Pico Water Systems

Complete all-in-one water system which provides Type 2 or Type 3 water with many unique features. It has low energy consumption, minimal water wastage and maximises the use of your laboratory space.

Model Specifications

| Model | RO Make-up Rate (@ 15°C) | Inorganics Rejection Rate | Organics Rejection Rate | Bacteria Rejection Rate | PH | Daily Usage (Max) |
| --- | --- | --- | --- | --- | --- | --- |
| PICO10T3 | 10lph | 98% | >99% | 99% | Effectively Neutral | 60lpd |
| PICO20T3 | 20lph | 98% | >99% | 99% | Effectively Neutral | 120lpd |
| PICO10T2 | 10lph | 1-15MΩ-cm | >99% | 99% | Effectively Neutral | 60lpd |
| PICO20T2 | 20lph | 1-15MΩ-cm | >99% | 99% | Effectively Neutral | 120lpd |

- Tap to Type 2 or 3 Water

- Production Rate of 10lph or 20lph

- Low cost, one box solution

- Small footprint

- Energy consumption <60w*

- Touch-screen interface

- Unique disposable RO module & pre-treatment cartridge

- On-bench, or wall mounting option available

- Built-in 35 litre reservoir

- Minimal water wastage

- Simple cartridge replacement

- First consumables included with unit

- Mercury-free UV option available

Dimensions

| Width | Height | Depth |
| --- | --- | --- |
| 485mm | 500mm | 330mm |

Weight

Weight (Empty)

20kg

Energy Consumption

| Model | Minimum Consumption* |
| --- | --- |
| PICO10T3 | 0.028kW |
| PICO20T3 | 0.039kW |
| PICO10T2 | 0.028kW |
| PICO20T2 | 0.039kW |

*dependent on feed water pressure

Services required

- 3/4” BSP isolated cold water feed (washing machine tap/valve)

- 240 Volt, 13 Amp sockets

- Drain/up-stand with trap or sink/tundish

Feed Water Requirements

| Pressure | Bar | – 5 |
| --- | --- | --- |
| Conductivity | µS-cm | <1400 |
| Colloid Index | SDI | <3 |
| Free-chlorine and Iron | Mg/l | 0.1 |